Operating the Hermes microscope

Set-up

- Open the Hermes software from the desktop (Hermes2537)

- Turn on the Hermes using the software control panel

  - All module indicators should turn green, this may take several minutes

  - “Close” the control panel once all module indicators are green

- Click on “noursadeghi lab”

Putting plate in machine

- Clean bottom of plate with Kimtech wipe and ethanol spray

  - Handle plate by the edges following this

- Open door using button on top bar of screen

  - The icon is the third one in from the right

- Place plate on the stage gently

  - It should rest flat on the supports

  - Push to back left corner of stage

- Close door using the “Close Door” onscreen prompt

  - NB all control of Hermes is via the computer- don’t try and manually close door or manually move any of the equipment

Selecting a protocol

- Click “protocol” tab in top bar

- Select “browse”

- Navigate to an image acquisition file with similar set up

  - Data/Hermes/Useres/noursadeghi lab/desired expt folder

- Open the sub-folder ending “IA_date”

  - It should contain a folder called “Reports”. Once this is visible- just click “select folder” on the dialogue box

- Protocol will appear in list on screen- double click to open

Setting up imaging and checking parameters

- Scan Set up

  - Select “Scan Setup”

  - Highlight wells of plate for imaging

    - Hold “Ctrl” to select multiple areas

  - Confirm selection with “select scan area”

  - You shouldn’t need to alter other parameters

  - Click “Done”

- “Imaging parameters”

  - Select “Image Parameters calibration”

  - Turn off the video

    - The default is for the light source and vide to be on- this will bleach the plate. Turn it odd by clicking the “video” button once the image parameter window opens

  - Navigating the image parameter set up

    - Toggle between channels using the thumbnails

    - Click on wells to view different wells

    - The joystick can be used to move within the well

      - Use it gently with small movements

      - Pull the trigger on the joystick to take a snapshot of the new position in the well (if the video link is off)

  - Optimise focus

    - If you have loaded a previous protocol this should be about right already

    - Check the blue and green channels and the wells at the corners of your imaging areas i.e. the extremes

    - If you need to edit the focus

      - Turn the video on (Video button)- try to keep this brief to avoid bleaching

      - Adjust offset

        - Coase focus: buttons 3/4 on diagram

        - Fine focus: buttons 5/6 on diagram

      - Set new focus button 2 on diagram

      - Turn video off

  - Optimise intensity/exposure/gain

    - Aim is to have a few saturated pixels within well so full dynamic range is being used

      - Too much saturation makes segmentation difficult during analysis

    - Tend to alter intensity and exposure. Keep gain 10-15%

      - Can alter settings using buttons in panel

      - Note, with the video off the image will not update in real time. To see the effect of any changes take a snap shot of the FoV using the trigger on the joystick

  - When compelte: click “Done”

- Click “Next”

  - Enter a name for the experiment output folder

- Click “Start”

At the end of imaging

- Check all wells have imaged (they should be green on the plate diagram)

- Remove plate

  - Open door using button on top bar of screen

    - The icon is the third one in from the right

  - Remove the plate and wrap in foil

    - Parafilm plate prior to storing in fridge

  - Close door using the “Close Door” onscreen prompt

Turning off the Hermes

- Use the “Exit” button

  - Icon is the second one in from the right on top bar

- You will return to initial screen

- Click the icon with the Hermes + green tick in the top right corner of screen (second in from right)

- The control panel will open

  - Click “off”

  - The module lights should turn from green to red and the Hermes should become silent

    - If either of these don’t happen let me know

  - Close the control window

- Exit the software using the “Power” icon in top right hand corner of screen

  - It will ask “are you sure you want to exit the program”-> select “Yes”

Cope imaging files to external HD

- Folder located in Data/Hermes/Useres/noursadeghi lab/desired expt folder